## Slide 1
Image Formation of Optical Coherence Tomography
Computational Optics Group, University of Tsukuba
Yoshiaki YASUNO
Disclosure
Nikon (F), Yokogawa (F), Topcon (F), Sky Technology (F), Kao (F), Sigmal LIS (F)

## Slide 2
Outline
What is OCT: history and sub-types.
Basic OCT-imaging property.
How to improve imaging performance.
Extension of OCT imaging function.
Why we need to revisit OCT imaging theory?

## Slide 3
Huang, Science (1991)
Huang et al., Science 254, 1178-1181 (1991).
Ex vivo porcine retina/optic nerve head
D. Huang, E. Swanson, C.P. Lin, J.S. Shuman, W. G. Stinson, W. Chang, M.R. Hee, T. Flotte, K. Gregory, C.A. Puliafito, J.G. Fujimoto (Medical, Engineering, MD-PhD)
Huang
Schuman
Fujimoto
Notes: http://www.octnews.org/topic/profile/david-huang/
http://www.mirm.pitt.edu/people/bios/Schuman1.asp

## Slide 4
Optical Coherence Tomography (OCT)
Point detection
2-D scan for 2-D image
long measurement time
Time domain OCT
High optical efficiency
Spectral domain OCT
with spectrometer
Swept source OCT
with wavelength-sweeping light source
Suitable for λ ≥ 1 μm
Line detection
1-D scan for 2-D image
Fourier domain OCT
Sub-types of OCT
Relatively low sensitivity
High-Speed and
3-D Tomography
Better S/N around +20 dB
Suitable for λ < 1 μm

## Slide 5
Basic Imaging Theory of OCT

## Slide 6
Structure resolving principle
3D resolution
Axial
Lateral
Low coherence
interferometry
Time-delay (Time-domain OCT)
Spectral interferometry (Fourier-domain OCT)
Point-scanning
= Reflection confocal
Full-field imaging
Spatially incoherent
Spatially coherent
Line-field imaging

## Slide 7
Structure resolving principle
TD-OCT
FD-OCT
Point-scanning
Full-field imaging
Line-field imaging
Classic TD-OCT
Huang, Science 254, 1178-1181 (1991).
Häusler, J. Biomed. Opt. 3, 21-31 (1998).
Classic FD-OCT
Current standard
Incoherent
Coherent
Dubois, Appl. Opt. 41, 805- (2002).
Hillmann, Opt. Express 25, 27770- (2017).
FF-SS-OCT
FF-OCT
NA
Nakamura, Opt. Express 15, 7103- (2007).
Fechting, BOE, 6, 2156- (2015).
LF-SD-OCT
Standard
High speed
High resolution

## Slide 8
Time-domain OCT

## Slide 9
Broadband Source
Line sensor
B-Scan
Low coherence interferometer
Spectrometer
Reference (Fixed)
Signal Processing
Coupler
Fourier-domain (spectral-domain) OCT

## Slide 10
Fourier Trans.
Spectrometer
Spectrum
Intensity
CCD
Power spectrum
k
Intensity
i-Fourier Trans.
DSP
Wiener-Khinchine's
Loop
Auto correlation
Conjugate
Cross-correlation
＝ OCT signal
Correlation distance
Correlation
Auto correlation
Time signal
Time = depth
Reference
Probe
Intensity
Principle of FD-OCT

## Slide 11
Resolutions of OCT
Axial resolution
= temporal coherence length
Lateral resolution
Scanning
Flood-illumination (spatially incoherent)
Examples
for λc = 840 nm, Δλ = 50 nm, n = 1.38
Δz = 4.5 μm in tissue
Open Issues
x-z separability assumption
Limited depth-of-focus
Speckle-limit of practical resolution
for f = 30 mm, d = 2 mm, point scan
Δx = 16.0 μm in tissue
for NA = 0.06, flood-illumination
Δx = 8.54 μm
Refractive index
Focal length
Beam width

## Slide 12
Axial-Lateral interaction
Villiger, JOSA A 27, 2216- (2010).
NA = 0.5
For high NA, the PSF skews in axial (z) - lateral (r) space!
Axial and lateral PSFs are not separable for high NA!

## Slide 13
Speckle in OCT
Coherent imaging
Inevitable speckle
Speckle size = resolution
→ limited structural visibility
Speckle reduction
Averaging
Sophisticated statistical image processing
Exploiting speckle
E.g., Schumitt, J. Biomed. Opt. 4, 95- (1995).
Spaide, Am. J. Ophthalmol 146, 496- (2008).
For example, Ralston et al.
DOI: 10.1007/978-3-319- 06419-2_14 (2015).
Z. Jian, Opt. Lett. 34, 1516- (2009).
OCT angiography
Mariampillai, Opt. Lett. 33, 1530- (2008).
Li, Opt. Lett. 44, 787- (2019).
Elastography
Incoherent FF-OCT
No speckle
E.g., Latrive, BOE 2, 2897- (2011)
Practical resolution cannot be described
by standard theory!

## Slide 14
High-Resolution Imaging

## Slide 15
Ultra-high-resolution OCT
Broad band light source
Visible Light OCT
Δz ~ 1 μm in tissue KLM Ti:Sa laser, TD-OCT 350-nm width @ 800 nm
Drexler Opt. Lett.
24, 1221- (1999).
Δz ~ 2.1 μm Ti:Sa laser, SD-OCT 144-nm width @ 800 nm
Wojtkoski, Opt. Express 12, 2404- (2004).
Northwestern Univ. Δz ~ 0.97 μm in air Super continuum, SD-OCT 506 - 621 nm
S. Chen, BOE 8, 1416- (2017).
In vivo human
In vivo human
Δz < 2 μm in air Super continuum, SD-OCT 100-nm width @ 575 nm
UC Davis
SP. Chon, BOE 8, 323- (2017).
Notes: SuperK EXTREME EXW-6, NKT Photonics

## Slide 16
High lateral resolution and long DOF
Computational/holographic refocusing
High-order aberrations
Forward-model based method
ISAM
(inverse scattering)
Numerical back-propagation
Sample plane
Numerical correction lens
Focal plane
Yasuno, Opt. Express, 14, 1006- (2006).
Oikawa, BOE 11, 5886- (2020).
Original
Refocused
(fx, fy, k)
(x, y, τ) -space
(fx, fy, fz)
Data remapping
(x, y, z)
Ralston, Nat. Phys. 3, 129- (2007). E.g., Coquz, OpEx 25, 30807- (2017).
Kumar, OpEx 21, 10850- (2013).
Adi, PNAS 109, 7175- (2012).
Ruiz-Lopera, Opt. Lett 45, 5982- (2020).
For phase-unstable system
Numerical volumetric phase stabilization

## Slide 17
Hidden-Assumption and Limitation
Signal in 4D frequency
τ-z difference
z
τ
fz
fx, fy
ω
fx, fy
fx, fy
ω
fz
Fourier Transform
Aperture
Fukutake, Photonics West/BiOS (2020).
OCT image spectrum
True sample structure spectrum
c
Integration
Integration
becomes thick
for high NA.
Conventional computational refocusing does not work for high NA!

## Slide 18
Multi-functional imaging

## Slide 19
Multi-functional imaging
Polarization
Temporal property
OCT
OCT angiography
Dynamics imaging
Polarization-sensitive OCT
Jones matric OCT etc…
Argumentation
Generalization

## Slide 20
Multi-functional imaging
Polarization
Temporal property
OCT
OCT angiography
Dynamics imaging
Polarization-sensitive OCT
Argumentation

## Slide 21
Polarization sensitive OCT
Phase retardation imaging = probe beam polarization
Birefringence / local axis imaging = tissue property (?)
Open Issue: Non-polarization polarization property imaging (?)
De Boer, Opt. Lett. 22, 934- (1997).
Gesperger, Appl. Sci. 9, 2100 (2019). Amyloid beta plaque
Sugiyama, BOE 6, 4951- (2015). Birefringence
Villiger, Optica 5, 1329- (2018). Local optic axis

## Slide 22
OCT angiography / flow imaging
Doppler OCT
Time-variance imaging (OCTA)
Z.P. Chen, Opt. Lett. 22, 1119- (1997). 1st Doppler OCT.
Baumann, BOE 2, 1539- (2011). En-face Doppler.
Makita, Opt. Express
14, 7821- (2006). Doppler OCTA
Y.L. Jia, BOE
20, 4710- (2012). SSADA
Plonar, Retina
36, S118- (2016). VISTA
D.Y. Kim PNAS
110, 14354- (2013). Phase variance

## Slide 23
Dynamics imaging
Apelian et al., Biomed. Opt. Express, 7, 1511– (2016).
Rat liver
I. A. El-Sadek et al., Biomed. Opt. Express, 11, 6231–(2020).
Tumor Spheroid
Münter et al., Opt. Lett., 45, 4766- (2020).
Murine Tongue
Leung et al., Biomed. Opt. Express, 11, 2768– (2020).
Human esophageal biopsy
Slide summarized by Pradipta Mukherjee
How do we describe “temporal imaging property”?

## Slide 24
Background and current issue: Why should we revise OCT imaging theory?
Label-free imaging
Deep-tissue imaging
Axial & Lateral high-resolution imaging
Thick tissue culture
(spheroid / organoid)
Regenerative medicine
Effective animal-study
Drug discovery
Holographic refocusing/AO
High NA
Axial-lateral interaction
OCT imaging theory to be revised.
HOT TOPIC!
OCT microscope
Multiple scattering
and Speckle